## Slide 1: XMM-Newton Calibration Updates
Michael Smith, ESAC
(Ivan Valtchanov, XMM Cal team)
15th IACHEC Meeting, 24 April 2023

## Slide 2: XMM-Newton calibration releases over 2022/23
Astrometry: Time variable boresight - 2023 update XMM-CCF-REL-394
RGS: Update of the correction to the RGS Effective Area XMM-CCF-REL-395
EPIC-pn: Spatial CTI correction and energy scale in [6,9] keV for XMM-Newton
EPIC-pn in Full Frame and Extended Full Frame modes XMM-CCF-REL-391
        Dedicated talk by I. Valtchanov in Detectors and Background WG session
EPIC-MOS: EPIC MOS contamination XMM-CCF-REL-390
EPIC MOS response XMM-CCF-REL-396
EPIC pn & MOS: Empirical correction of the EPIC effective area XMM-CCF-REL-388
Dedicated talk by Felix Fürst (later today)
Notes: Astrometry and RGS eff area: annual/periodic updates using data up to the end of 2022.

## Slide 3: EPIC-MOS: contamination and redistribution
MOS cameras show time dependent changes in response due to
contamination (Aeff change < ~ 1 keV)
spatially dependent redistribution
  patch core: r<14”
  patch wings: 14<r<36”
  off-patch: r> 40”
  where main photo peak “shoulder” flattens into a “shelf” in patch
Are corrected for in calibration (contaminant model; epochal & spatially dependent RMFs) but need periodic updates. Previous major updates were in 2013.
Degeneracy: a given RMF solution is dependent on the Aeff.
Current methodology:
Update contaminant model (off-patch data from SNR 1E 0102)
Update RMF based on new Aeff estimate
Notes: How the new Aeff is estimated? 1: contamination updates Aeff and then with the Aeff the RMF si updated.

## Slide 4: EPIC-MOS: contamination
Measurement of contaminant based on off-patch observations of SNR 1E0102-7219
Modelled as pure C
Depth determined by the ratio of observed count rates in 0.1 - 0.75 keV and 0.98 - 3.0 keV bands
Depth as function of time modelled by an exponential (MOS2)
Ratio of the MOS2 effective area curves from
revolutions 3652 (~2019) and 1082 (~2005)
S.Rosen
Notes: Why there is a gap from 2008 to 2013? To ask Simon/Jari

## Slide 5
EPIC-MOS: redistribution
Redistribution:
Last significant update was in 2013
New release is a substantial extension of the time-base: 5 new epochs of 300 revs each, from 2013 onwards
Built upon the latest update of MOS time-dependent contamination
Based on iterative adjustment of empirical RMF model, given data and assumed spectral models of
  1E 0102-7219
  Zeta Puppis
  RX J1856.5-3754
Sources are observed on- and off-patch to allow
RMFs to be produced for on-patch, patch wings and off-patch regions

## Slide 6
Previous broad temporal behaviour continues
Especially for the core and wings regions:
Drift to lower energies of the peak of the low-energy shoulder component
Broadening of its low energy wing
EPIC-MOS: redistribution
Epochal evolution of redistribution functions @ 300 eV
S.Rosen
Notes: Colour coded are the epochs. More important for core and wings and less so for off patch

## Slide 7
EPIC-MOS: contamination and redistribution
Results tested on:
RX J1856 and N 132D: spectral fits (using IACHEC models, free overall normalisation) show significant improvement in fit quality.
RX J1856: temporal flux stability (in 0.2-0.5 keV band) is maintained.
RX J0720.4-3125: (NS, observed in rev 3636) is a case showing worsening results – could be due to uncertainties in the model used.
MOS1
MOS2
old calibration
N 132D
new calibration
S.Rosen
Notes: Temporal flux stability: is this something to check on my side for the SNR flux stability issues?

## Slide 8
EPIC-MOS: contamination and redistribution
Results tested on:
RX J1856 and N 132D: spectral fits (using IACHEC models, free overall normalisation) show significant improvement in fit quality.
RX J1856: temporal flux stability (in 0.2-0.5 keV band) is maintained.
RX J0720.4-3125: (NS, observed in rev 36363) is a case showing worsening results – could be due to uncertainties in the model used.
Sample of 37 sources: stacked data/model ratios used to evaluate MOS-to-PN cross calibration; improved consistency between MOS1 and PN at lowest energies.
old calibration
new calibration
Stacked data/model, normalised to PN

## Slide 9: RGS: contamination and effective area
Extra contamination of
undetermined origin
Contamination estimated from the flux of the Isolated Neutron Star RX J1856-357 at long wavelengths.
Observed flux decrease not explained by C8H8 contamination.
Empirical corrections in place – latest update covers data taken until end of 2022.
R. González-Riestra
Notes: The curve is the previous model?

## Slide 10: RGS: wavelength scale
Average wavelength shift of spectra of spectra of bright stars:
Measured shifts of emission lines in two samples of stars: “Bright” (orange) and “Faint” (cyan):
R. González-Riestra
Notes: So, what’s the conclusion? No problems here? Within 5 mA

## Slide 11: OM: time-dependent sensitivity degradation
S.Rosen
OM throughput estimate for 2030:
Time-dependent degradation continues to be monitored, and is corrected in the SAS
Notes: Degradation to what? To magnitudes/flux, sensitivity, throughput?

## Slide 12: The end