## Slide 1
NICER Calibration Status Craig Markwardt (NASA/GSFC) on behalf of NICER Team
GSFC
IACHEC May 2024

## Slide 2: NICER’s Unique Capabilities
Launched June 2017 to International Space Station
Spectral band: 0.2–12 keV
  52 operating single-pixel silicon detectors
Energy resolution: < 150 eV @ 6 keV
  Comparable to X-ray CCDs
Timing resolution:
  100 nsec RMS absolute
Non-imaging field of view (“single-pixel”)
  6 arcmin diam. (half-max)
High throughput (3.5 Crabs with no pileup)

## Slide 3: Outline
NICER Calibration Status
Optical Light Leak and Repair
SCORPEON background model update
Future Calibration Plans

## Slide 4: NICER Calibration Status
NICER Crab
NICER calibration remains very stable
  No appreciable changes since 2021
Major impact of optical light leak
Other calibration aspects
  Energy scale: monitoring drifts of few eV / yr
  ARF: systematic feature errors <1.5% (0.4-10 keV)
  RMF: Scholze & Procop X-ray silicon model

## Slide 5: NICER Optical Light Leak
Onset
Impacts
Repair

## Slide 6: Light Leak Onset
Onset is apparent near 13:15 UTC on 2023-05-22
Increased levels of optical light loading (“undershoots”) during orbit day

## Slide 7: NICER X-ray Concentrator Assembly

## Slide 8: NICER Damage Due to Light Leak
Photographic survey shows damaged optical films
Higher fidelity surveys have been performed

## Slide 9: Light Leak Impacts
Primary impact is increased optical light loading
  Increased optical light by factor of 5x-10x during orbit day
  Much higher telemetry loads
  Extreme loading can push energy scale out of calibration
Consequences
  NICER avoids targets within ~100 deg of sun during orbit day
  Reduced ToO and long-term monitoring opportunities
  Lower level X-ray threshold increased to ~0.43 keV during orbit day to avoid noisy telemetry
X-ray performance is unchanged!
  Data during orbit night is still excellent, much of orbit day data is excellent as well, within observing constraints
  No detectable change to X-ray throughput for modules with known holes, to within few percent (Jeremy Hare work)

## Slide 10: Other Calibration/Software Impacts
Response and background calculators re-parameterized to handle day/night threshold adjustments
Better screening for science data

## Slide 11: NICER Repair is Approved
Repair patches installed during EVA (“space-walk”)
Jan – approved
May – flight hardware completed
May-Jul – testing
Aug 3 – planned launch
Nov – possible installation EVA
  Install locations based on photo survey (5 targets, 7 spares)
NASA Neutral Buoyancy Lab Testing

## Slide 12: Post-Repair Calibration Plans
Each patch removes exactly 1/6th of module effective area
”Delta” calibration plan
  Verification of successful patch by pointing near sun during orbit day, noting reduction in optical loading
  Observation of bright X-ray target (Crab, if visible) to verify effective area losses in repaired modules

## Slide 13: NICER SCORPEON
SCORPEON is a parameterized NICER background model
  Attempt at physically motivated components
  A-priori estimated norms, but XSPEC fitting allowed
Released Jan 2023
Feedback: SCORPEON becoming popular, primarily it always provides an estimate
  21 refereed pubs
SAA
Cosmic Rays (COR)
Precipitating Electrons (PREL)
Trapped Electrons
(TREL)
Thorne et al 1980 Tyssoy presentation

## Slide 14: Summer 2024 Calibration Updates
Several known minor problems in response generation
  Expect ~5% change in flux scale
    Based on improved deadtime understanding
  Expect 0.05 change in power law shape
    Remove non-physical fit of optic geometric area
  Expect low energy threshold shift ~5 eV
    Based on observations of HZ43, RX J1856, Crab
  Deal with NICER patch repair (loss of area due to patch blockage, 5-7 pie wedges)
The goal will be to maintain “old approach” and “new approach” calibration in parallel for some time