## Slide 1
2026-4-22
The Spectroscopy Focusing Array Timing (SFA-T) payload onboard the eXTP mission：development status, preliminary ground calibration results and future calibration roadmap
18th IACHEC meeting, Pelham, Germany, April 20-24 2026
Yuxuan Zhu
zhuyx@ihep.ac.cn
IHEP, CAS
On behalf of the eXTP SFA-T Team
YuSa Wang1, Wei Li1, ZhenJie Li1, LaiDan Luo1, Jia Ma1, Qiong Wu1, ZiJian Zhao1, XiongTao Yang1, YuXuan Zhu1, YiFan Zhang1, Can Chen1, ZiYu Bao1, YaoGuang Liu1, Wei Xu1, Dong Xie1, XueZhao Li1, JunYi Tian1, YiChen Lin1, Fan Yang1, XueYi Tang1, Gang Li1, Juan Zhang1, LiQiang Qi1 , TianXiang Chen1, MaoShun Li1, Bin Meng1, HongWei Liu1, XiaoJing Liu1, YanJi Yang1, Juan Wang1, XiaoFan Zhao1, WeiWei Cui1, Yong Chen1, YuPeng Xu1, FangJun Lu1, ShuangNan Zhang1, HuiLin He1, Hua Feng1, CongZhan Liu1, Bo Wang2, LangPing Wang2, DianLong Wang2, Fei Ding2, LiZhi Sheng3, PengFei Qiang3, Norbert Meidinger4,Thomas Bechteler4, Vadim Burwitz4, Kirpal Nandra4, Stefano Basso5, Giorgia Sironi5, Giovanni Pareschi5, Tagliaferri Gianpiero5. （Updated Aug.2025）
1. Institute of High Energy Physics, CAS, Beijing, China; 4. Max-Planck Institute for Extraterrestrial Physics, Giessenbach Strasse 1, 85748 Garching, Germany;
2. Harbin Institute of Technology, Harbin, China; 5. INFN-Brera Astronomical Observatory, Via E. Bianchi 46, 23807 Merate, Italy;
3. Xi’an Institute of Optics and Precision Mechanics, CAS, Xi’an, China;

## Slide 2: Outline
Configuration of eXTP and SFA-T
1
Development status of SFA -T
2
Preliminary calibration results
3
Ground calibration strategy for future
4
Summary
5

## Slide 3: Outline
Configuration of eXTP and SFA-T
1
Development status of SFA -T
2
Preliminary calibration results
3
Ground calibration strategy for future
4
Summary
5

## Slide 4: 1. The configuration of eXTP
What are building？ eXTP: enhanced X-ray Timing and Polarimetry Mission
Payload | Configuration | Optics | Detector | Eff. area (m²)
Spectroscopy Focusing Array (SFA) | 9 telescopes | Wolter-I, nickel | SDD, 0.5-10 keV 180 eV @ 6 keV | 0.6 m²(1-2 keV)
Large Area Detector (LAD) | 40 modules | MCP | SDD, 2-30 keV 180 eV @ 6 keV | 3.1 m² (2-10 keV)
Polarimetry Focusing Array (PFA) | 4 telescopes | Wolter-I, nickel | GPD, 2-8 keV 1.8 keV @ 6 keV | 380 cm² @ 3 keV
Wide Field Monitor (WFM) | 6 cameras | Coded mask | SDD, 2-50 keV 500 eV @ 6 keV | 3 Sr (FOV)
Scientific goal: eXTP is a science mission designed to study the state of matter under extreme conditions of density, gravity and magnetism.
Operation：Pointing observation with high accuracy and stability.
Configuration before 2024

## Slide 5: The configuration of eXTP
Payload | Configuration | Optics | Detector | Eff. area (cm²)
Spectroscopy Focusing Array (SFA) | SFA-T 5 telescopes | Wolter-I, nickel | SDD, 0.5-10 keV 180 eV @ 6 keV | 3300 cm² @1.5 keV 2000 cm² @6 keV
 | SFA-I 1 telescope | Wolter-I, nickel | pnCCD, 0.5-10 keV 180 eV @ 6 keV | 3.1 m² (2-10 keV)
Polarimetry Focusing Array (PFA) | 3 telescopes | Wolter-I, nickel | GPD, 2-8 keV 1.8 keV @ 6 keV | 180 cm² @ 3 keV
Scientific goal: eXTP is a science mission designed to study the state of matter under extreme conditions of density, gravity and magnetism.
Operation：Pointing observation with high accuracy and stability.
Orbit: Low Earth Orbit 610km 20°
Life: 5 years
Spectrum, Timing, Polarization, Imaging

## Slide 6: The configuration of eXTP
9 Mirror Modules
SFA-T *5
SDD
PFA *3
GPD
SFA-I *1
pnCCD
6 SFA & 3 PFA
🛰️ Satellite Platform • eXTP mission adopts satellite bus provided by Innovation Academy for Microsatellites, CAS, ensuring high reliability and streamlined integration.
🔍 Common Optics, Specialized Coatings • All three payloads share an identical geometric optical configuration; • Differentiation achieved solely through tailored reflective coatings, optimized for respective energy bands / scientific objectives.
X-ray

## Slide 7: The configuration of the SFA-T
 | SFA | PFA
Type | Wolter-I 47 shells 300mm + 300mm for P & H bolid | 
Focal Length | 5.25 ± 0.05 m | 
Coating | Au | Ni+Au
Effective Area | ≥820cm²@1.5keV ≥550cm²@6keV | ≥800cm²@3keV ≥550cm²@6keV
Energy Range | 0.3~10 keV | 
FOV | ≥ 12′ | 
HPD | ≤ 1′ (HPD), ≤ 3′ (W90) | ≤ 30″（HPD）
Profile | ≤ 600 mm（Diameter） | 
Weight | 150 kg | 
Temperature | 20±1℃ | 
The X-ray optics for eXTP

## Slide 8: The configuration of the SFA-T
SFA-T | Multi-Cell SDD
Number of cells | 19
Size | ≥ 18×18 mm2 (12′)
Energy range | 0.5-10 keV
Energy resolution | ≤180 eV@ 6 keV
Timing | ≤ 10 μs
SFA-T: 19-Cell SDD array
Filter Wheel
Flange
Shield
Electronics
SFA-T camera assembly
Six positions in FW：3 PI+Al filter、Fe-55、Close and Open.
SFA-I: PN-CCD
SFA-I | pnCCD
Pixel num | 384*384
size | ≥ 28.8×28.8 mm2 (12′)
Energy range | 0.5-10 keV
Energy resolution | ≤180eV@1.49 keV
Timing | ≤50 ms （FF） ≤ 240 μs （TM） ≤ 3 ms（PW）
SFA-I camera assembly

## Slide 9: The configuration of the SFA-T
The performance of SFA-T
Content | Requirements
Focal Length | 5.25±0.01m
Effective area （E±0.5keV） | ≥2750cm²@1.5keV ≥1670cm²@6 keV
Energy range | 0.5-10 keV
Energy resolution | ≤180eV@6 keV
FOV | ≥ φ12′
SDD pixel size | 6.4mm (4.2′)
HPD | ≤1′(HPD)，3′(W90)
Time resolution | ≤10μs
Absolute time accuracy | ≤ 2 μs
Dead time | ≤ 5% @ 1 Crab
Sensitivity | 4.5×10-15 erg cm-2 s-1 (0.5-10 keV, 5σ, 106 s)
Content | Requirements
Focal Length | 5.25±0.01m
Effective area （E±0.5keV） | ≥550cm²@1.5keV ≥330cm²@6 keV
Energy range | 0.5-10keV
Energy resolution | ≤180eV@1.49 keV
FOV | 12′ ×12′
HPD | ≤1′(HPD)，3′(W90)
Time resolution | ≤50 ms （Full frame） ≤ 240 μs （Timing mode） ≤ 3 ms（Window mode）
Sensitivity | 4×10-15 erg cm-2 s-1 (0.5-10 keV, 5σ, 106 s)
The performance of SFA-I
SFA-T (SDDs): Optimized for timing performance — high time resolution
SFA-I (pnCCD): Optimized for imaging capability

## Slide 10: Outline
Configuration of eXTP and SFA-T
1
Development status of SFA -T
2
Preliminary calibration results
3
Ground calibration strategy for future
4
Summary
5

## Slide 11: Development status
Who is building all this？
Lead Institution
• IHEP, CAS (China) — Overall project coordination, system integration & technical leadership
International Partners
• MPE, Germany — Provision of SDDs
• INAF, Italy — Technical support for optical component design
Chinese Consortium Responsibilities
✅ Electronics & Data Acquisition
✅ Thermal Control & Mechanical Structure
✅ Optical System Integration：Jointly developed by IHEP, HIT and XIOPM
✅ End-to-end Testing & Calibration

## Slide 12: Manpower and resources
20+ experts for the payload development！
Dedicated Detector Team: Professional team specializing in detector design, fabrication, and end-to-end testing.
Calibration Capability: Payload team operates China's largest ground-based calibration facility; tailored calibration plan developed for SFA-T technical requirements.
Industry Partnership for Chip : Technical support secured from leading large-format semiconductor manufacturers, enabling full-cycle chip design and validation.
Cross-Calibration Access: Leveraging PANTER facility support for high-precision cross-calibration between optical system and detector subsystems.

## Slide 13: Calibration facilities
The biggest X-ray test facility in China
Calibrated the EP FXT and WXT well
X-ray
sources
Pump stations
Big Chamber
100XF (100-m X-ray Test Facility):
  Mirror calibration (EA, PSF, Time resolution)
  Telescope calibration
PANTER Facility:
  Cross calibration of mirrors

## Slide 14: Calibration facilities
Individual chamber for the SDDs test and calibration
  Electronic development
  Time resolution
  Energy scale
  Edge block
  …
synchrotron radiation source:
  RMF
  QE
  Flux limitation
  …

## Slide 15: Calibration facilities
A big housing system is developed for the telescope calibration
  Material：Ti-TA1
  Load capacity: 850 kg
  Angular positioning accuracy: ≤5 ″
  Attitude Monitoring System with a position sensitivity of 10 μm（0.5″）
  Deformation variation with the angle changing (±1°)：0.04 mm(2″)
synchrotron radiation Cal chamber
  Equipped with a ion pump assembly, achieving ultra-high vacuum
  Integrated beam flux monitoring provides live feedback on detector exposure
  Upon safety thresholds, instantly close the valve, shielding the detector from beam damage.

## Slide 16: Outline
Configuration of eXTP and SFA-T
1
Development status of SFA -T
2
Preliminary calibration results
3
Ground Calibration strategy for future
4
Summary
5

## Slide 17: Preliminary cal. results
Many individual shells had been tested since the Phase A, and HPD ~26″ can be achieved!

## Slide 18: Preliminary cal. results
Mirror module#3 (MM3)had been cross-calibrated at both 100XF and PNATER
Calibration in IHEP, Beijing
2024.5
Calibration in PANTER, Munich
2024.8

## Slide 19: Preliminary cal. results
Calibration in IHEP, Beijing
2025.9
Energy/keV | HPD/″ | W90/″
0.28 | 26.34±0.45 | 113.01±1.62
0.93 | 27.17±0.33 | 119.51±1.22
1.49 | 27.36±0.34 | 125.51±1.32
2.98 | 28.78±0.32 | 150.24±1.40
4.50 | 33.01±0.38 | 178.41±1.73
5.40 | 40.38±0.56 | 205.26±2.39
6.40 | 43.40±0.69 | 197.93±2.64
8.04 | 41.22±0.67 | 155.07±2.11
An Engineering module of PFA were also tested at 100XF
A bit of worse than the MM3, but still meets the requirements of SFA-T

## Slide 20: Preliminary cal. results
SDDs manufactured in IHEP is under calibration!
Even a chip with relative poor energy resolution was chosen, it still satisfies the energy resolution requirements of SFA-T
A capability of pulse profile observation
Time resolution remains to be verified!

## Slide 21: Preliminary cal. results
Pixel Edge Effect
When pixel edges are physically masked, the measured FWHM appears artificially improved (195 eV → 167 eV @ 5.9 keV);
Time-window-based event reconstruction fails for interactions near pixel edges;. There are some charge lost during the charge transfer!
A mask (edge block) will be installed above the SDDs chip.

## Slide 22: Preliminary cal. results
Conservative Design (1 mm thickness wall)
PSF-convolved simulations show ≤3.5% eff. Area loss;
Optimized Solution (300 μm thickness wall)
Further analysis confirms 300 μm is sufficient to block invalid edge events;
Resulting effective area loss: only ~0.8% — statistically negligible for scientific performance.

## Slide 23: Preliminary cal. results
150eV@6 keV
The multi-cell SDD is being developed well in MPE!

## Slide 24: Outline
Configuration of eXTP and SFA-T
1
Development status of SFA -T
2
Preliminary calibration results
3
Ground calibration strategy for future
4
Summary
5

## Slide 25: Ground calibration strategy for future
The calibration content is listed, huge calibration work, also test in the next 5 years.
Device | Items
Mirrors level | Effective area (E,φ)
 | PSF(E,φ)
 | FoV
Detector level | Edge effect
 | Energy threshold
 | QE(E)
 | Energy resolution, RMF(E)
 | Time resolution
 | Readout noise
Telescope level | ARF(Effe. Area of mirror, filter transmission, detector QE,etc)
 | Vignetting considering the edgeblock
Full telescope-level calibration cross-checks mirror + detector sub-system results, guaranteeing end-to-end data consistency.
High-volume calibration workload for 5 telescopes;
Beam divergence significantly biases mirror calibration (esp. eff. area);
Time constraints preclude fine-sector scans divergence correction.
Hybrid approach: Simulation model fits beam calibration data to extract physical parameters (mirror roughness, figure error);
Derived parameters feed into the CalDB, ensuring traceability and credibility.

## Slide 26: Ground calibration strategy for future
End of 2026
March, 2027
QM cal.(optics, detector, telescope):
Facility & Equipment Qualification will be test
Calibration Strategy Validation & Iteration
FM cal.(optics, detector, telescope) start
Adopt a "build-then-calibrate" pipeline: each FM telescope undergoes calibration immediately after integration;
Resource Estimate: 530 workdays for Calibration
Mid of 2027
2030
Launch!
Cross calibration Maybe!

## Slide 27: Outline
Configuration of eXTP and SFA-T
1
Development status of SFA -T
2
Preliminary calibration results
3
Ground calibration strategy for future
4
Summary
5

## Slide 28: Summary
Experienced Team: Established a professional team with extensive development expertise from prior missions, including EP-FXT and HXMT.
Phase B Initiated: Officially entered Phase B; active QM development is underway for both cameras and optics.
Robust Calibration Facilities: Existing facilities (100XF, PNATER) fully meet requirements; other equipments are under development.
Comprehensive Test Schedule: Extensive testing and calibration campaigns will be conducted across 2026–2027, with an estimated total active duration of 2–3 months.
Clear Path to Flight Model: High confidence in successful QM completion, ensuring a seamless transition to final FM development.

## Slide 29
Thank you!